The STEM Digital Twin

A guide to what it simulates and how to drive it

# 1. What it is

A simulated scanning transmission electron microscope with a vendor-neutral control interface. It renders images and electron-diffraction patterns from thirteen specimens, models the instrument behaviours that constrain a real acquisition — finite depth of field, stage drift, dose-driven contamination, stage safety limits, autofocus that can decline to converge — and exposes all of it through an interface whose every method has a counterpart on a real column.

Diffraction is computed kinematically. That is a deliberate choice of operating point: the engine runs in seconds on a laptop processor with no specialised libraries, which is what makes it usable inside an interactive acquisition loop. Dynamical diffraction is available offline through an appendix, on the same specimens.

The purpose is to let acquisition and decision logic be written, run and debugged without an instrument, and then moved onto one by exchanging a single object.

# 2. How it is put together

A server holds the simulation and answers remote procedure calls over a socket. Clients connect to it. The server can run inside the same session as the client or as a separate process; either works.

Two client objects, and the split between them is the point:

| Object | What it is for |
| --- | --- |
| MicroscopeControlClient | Everything a real instrument can also do: move the stage, set the beam, choose imaging or diffraction, set magnification or field of view, acquire, autofocus, set optics. |
| SimulationHarness | Everything only a simulation can do: choose the specimen, set its thickness, inject drift and contamination, reset the accumulated dose map. |

A workflow written against the control client alone carries no dependency on the simulation. That is what makes the boundary between testing and deployment explicit rather than assumed.

# 3. The specimen

## 3.1 Geometry

Every specimen is generated within a volume spanning 20 µm laterally and 100 nm in depth. The working thickness and its position within that depth are chosen when the sample is loaded, so a session can be run on a thin region or a thick one. Thickness enters the image through mass-thickness contrast and enters diffraction through the width of the relrod.

Beyond that, the specimens differ in their lateral extent. The single-crystal and dislocation models are FIB-lamella-like: a strip approximately 5 µm wide with vacuum beyond its edges, so driving the stage across it leaves the specimen, as it would on a real grid. The remaining models fill the field and have no edge to find.

## 3.2 The samples

| Identifier | Specimen | Class | Finest feature (nm) |
| --- | --- | --- | --- |
| fcc_single_crystal | Fe (FCC, γ-austenite) | Single crystal | 0.20 |
| bcc_single_crystal | Fe (BCC, α-ferrite) | Single crystal | 0.20 |
| hcp_single_crystal | Mg (HCP) | Single crystal | 0.20 |
| polycrystal_grains | Polycrystal (Fe FCC, few grains) | Multi-grain | 0.25 |
| dislocation_crystal | Fe FCC [111] with edge dislocations | Defected crystal | 0.25 |
| amorphous_film | Amorphous film | Amorphous | 0.30 |
| au_dispersed | Au nanoparticles, dispersed | Nanoparticle | 2.0 |
| au_clustered | Au nanoparticles, clustered | Nanoparticle | 2.0 |
| au_bimodal | Au nanoparticles, bimodal size | Nanoparticle | 1.5 |
| au_on_substrate | Au on substrate (supported catalyst) | Nanoparticle | 2.0 |
| core_shell | Core–shell nanoparticles | Nanoparticle | 1.0 |
| shape_assembly | Synthetic convex features | Feature testbed | 30.0 |
| atomsk_polycrystal | File-driven (Atomsk / LAMMPS / XYZ / CIF) | External structure | — |

The last column matters for magnification. Each sample declares the size of its finest real detail, and the renderer will not show detail below it. Resolving a 0.2 nm lattice therefore requires genuinely high magnification, at which point drift and dose become the limiting factors, exactly as on an instrument.

## 3.3 One definition of the specimen

Every sample implements a single method:

positions_A, Z = sample.get_atoms_in_region(cx_um, cy_um, half_width_um, depth_nm)

Anything needing the specimen’s atoms goes through it: diffraction, composition, and the offline multislice appendix. There is one source of truth for what a sample is, so no two parts of the system can disagree about it.

## 3.4 Roaming

The stage and accumulated drift both act in absolute coordinates, so a workflow driven back to a nominal position finds the field displaced by whatever drift occurred. Neither stops at a boundary.

Twelve of the thirteen samples treat their volume as one tile of a repeating specimen, with a period of 20 µm. The stage therefore never runs out of specimen, but structure recurs beyond that range: a scan wider than 20 µm, or a grid whose extent exceeds it, will revisit ground it has already covered. The synthetic-feature testbed generates from a position hash instead, so it never repeats — drive away and back and the same features are there, because they were never stored.

# 4. Forming an image

The render path, in order:

| Stage | What it does |
| --- | --- |
| Window | Absolute world position from stage plus accumulated drift. |
| Projection | Rigid-rotation tilt and per-depth-slice defocus, composited with coverage normalisation. |
| Thickness | Projected mass scaled by the working fraction of the specimen, then converted to a scattering fraction. |
| Probe | Point-spread function from aperture, spherical aberration and uniform defocus. |
| Resolution | Blur tied to the sample’s finest real feature, so detail appears only when the pixel is small enough. |
| Atomic columns | Crystalline samples at high magnification only. |
| Contamination | Additive brightening from the accumulated dose map. |
| Noise | Poisson dose statistics plus a Gaussian readout term. |

## 4.1 Stage tilt

Tilt is a rigid rotation of the specimen, about y by β and then about x by α, with the projection integrating along the laboratory beam axis:

x_lab = xs·cosβ − zs·sinβ
y_lab = ys·cosα − (xs·sinβ + zs·cosβ)·sinα
z_lab = ys·sinα + (xs·sinβ + zs·cosβ)·cosα

Rendering inverts the first two for each depth slice. Both stay affine in the output coordinates, so the cost is one resampling per slice. In-plane extent compresses by cosθ, which tracks the exact value to about one per cent across the tilt range.

## 4.2 Depth of field

The third line above is the distance along the beam, which means the specimen’s own depth, uniform defocus and the tilted focal plane are one expression:

dz_nm = nm_per_pixel · z_lab + defocus_nm

Blur width is proportional to that distance and clipped at a ceiling:

σ = |dz_nm| · (dof_max_sigma_px / dof_focus_gain_nm),  clipped at dof_max_sigma_px

dof_focus_gain_nm is the depth of field, expressed as the defocus that produces full blur. Its counterpart on a real column is the convergence semi-angle, through DOF ≈ 2λ/α²:

| dof_focus_gain_nm | Equivalent convergence semi-angle |
| --- | --- |
| 1000 nm | ≈ 2.2 mrad |
| 350 nm | ≈ 3.8 mrad |
| 60 nm | ≈ 9.1 mrad |
| 25 nm | ≈ 14.2 mrad |

It governs two things at once. Whether the specimen’s internal depth can matter at all — if the depth of field exceeds the foil thickness, everything is in focus wherever it sits — and the width of the in-focus band on a tilted specimen, which is dof_focus_gain_nm / (1000 · sin α) micrometres.

Depth slices are sorted into blur levels, each level blurred once and the levels summed with coverage normalisation. Because the mechanism reads only the voxel volume that every sample produces, all thirteen get depth of field with no specimen-specific physics.

## 4.3 Probe and resolution

A point-spread function carries the probe: aperture size, spherical aberration, and the broadening that uniform defocus causes. Its width is expressed in pixels, so it grows as the field narrows. A separate blur ties resolution to the sample’s declared finest feature.

# 5. Diffraction

Diffraction is computed from atomic coordinates in the illuminated region rather than from the image. For scattering vector q:

I(q) = | Σⱼ fⱼ(q) · exp(2πi q·rⱼ) |²

evaluated on the Ewald sphere, with atomic form factors and a thickness-dependent relrod — a sinc² in excitation error whose width scales as the reciprocal of thickness, so a thin region gives elongated reflections and a thick one gives sharp ones.

One code path serves every sample. Crystals give spot patterns, the polycrystal gives ring-like maxima, the amorphous film gives diffuse halos. The pattern depends on where the stage is: an aperture placed inside a single grain gives one net, and one straddling a boundary gives two superposed nets.

# 6. Detection and dose

Electrons per pixel follow from probe current and per-pixel dwell time:

nₑ = I · t / e

Counts are drawn from Poisson statistics on that number, scaled by detector quantum efficiency, with a Gaussian readout term added. Dose is therefore not a free parameter: shortening the dwell to limit exposure produces a correspondingly noisier image, and the trade-off a dose-limited workflow has to manage is present in the model rather than assumed.

The same dose figure drives contamination. Because dose is electrons per pixel divided by pixel area, it rises as the square of resolution: doubling the pixel count at a fixed field quadruples the dose delivered to the specimen. Carbon accumulates in the exposed region, and since the signal rises with mass-thickness that region brightens, leaving a footprint a later survey has to navigate around.

# 7. Instrument behaviour

## 7.1 Stage safety

Soft limits of ±1.5 mm in x and y, ±1 mm in z and ±30° in α and β. A move whose target lies outside any of them is rejected outright and nothing moves, as a real column would refuse it.

## 7.2 Magnification

Magnification is exposed on the control interface and on every backend, calibrated as mag = k / FOV. Setting either magnification or field of view keeps the other consistent.

## 7.3 Drift

Stage drift is specified in nanometres per second, covering the range from below 0.2 for an excellent column to roughly 10 for a just-inserted holder, with intra-frame shear and line-to-line jitter. It accrues over the real elapsed time between acquisitions, capped by max_dt_s so that an idle session does not displace the field on the next exposure.

## 7.4 Autofocus

The routine sweeps focus and assesses the sharpness curve for peak prominence above baseline and for multiple comparable maxima. It returns a convergence flag, a readable reason, the curve contrast and the number of candidate peaks. Where it does not converge it declines to commit and returns the stage to its starting height rather than guessing, so a workflow has to branch on the outcome. It is shown declining to converge on a heavily contaminated specimen, reporting insufficient peak-to-floor contrast as the reason.

# 8. Setting up an acquisition

Conditions are set by explicit calls, each stating a physical quantity. Most sit on the simulation harness; the autofocus criterion sits on the control client, because a real instrument has one too.

| Call | Sets |
| --- | --- |
| sim.set_drift(vx_nm_per_s, vy_nm_per_s, line_jitter_px, enabled, reset_accum, max_dt_s) | Stage drift in physical units. |
| sim.set_contamination(enabled, rate) | Carbon build-up. rate is a percentage: 100 is the nominal rate, 200 twice as fast, 0 off. |
| sim.set_noise(dwell_us, dqe, readout_e, use_dose_model, noise_sigma) | Dose and detector parameters. |
| sim.reset_specimen() | Clears the accumulated contamination map. |
| control.set_autofocus_limits(min_contrast) | Peak-to-floor ratio below which autofocus reports non-convergence. |

A worked example — a moderately drifting column with contamination running:

sim.set_drift(vx_nm_per_s=1.2, vy_nm_per_s=0.7, enabled=True)
sim.set_contamination(enabled=True, rate=100)
sim.set_noise(dwell_us=15.0, dqe=0.8, readout_e=1.5)
control.set_autofocus_limits(min_contrast=0.10)

Everything the acquisition sees is visible at the call site, so a demonstration states its conditions as numbers rather than referring to a named preset that has to be looked up.

# 9. Practical notes

Behaviours that are easy to misread as faults:

| Symptom | Explanation |
| --- | --- |
| Drift looks stationary | Drift is measured in nanometres per second, so whether it is visible depends on the size of an output pixel. At a 30 µm field on 512 pixels each pixel spans 59 nm, and 2 nm s⁻¹ needs about half a minute to move the image by one pixel. At 0.2 µm the same rate covers ten pixels in two seconds. Raise the magnification or wait longer; the rate itself is already realistic. |
| A long pause produces no drift | Drift accrues over the elapsed time between acquisitions, capped by max_dt_s so that leaving a session idle cannot teleport the field. The cap defaults to two seconds, so a deliberate hundred-second wait needs set_drift(max_dt_s=120) to take effect. |
| Contamination never appears, or appears instantly | It is observable over a window. Below roughly 0.3 e⁻ Å⁻² per frame nothing accumulates visibly; above roughly 20 the first frame already saturates. At 80 pA and 20 µs that window is about 5 to 30 µm of field at 1024 pixels, and it shifts down fourfold in field for every doubling of resolution. |
| The tilted in-focus band is invisible | Its half-width is an absolute distance, dof_focus_gain_nm / (1000 · sin α) micrometres, so what matters is its size as a fraction of the field. At a shallow depth of field and a wide field it is a narrow sliver. Narrow the field, or raise dof_focus_gain_nm if a wider band is wanted. |
| Defocus stops getting worse | Two ceilings apply. Depth-of-field blur saturates once the defocus reaches dof_focus_gain_nm, which is deliberate and tunable. The probe point-spread function is capped by its kernel radius, which binds sooner at higher magnification because its width is expressed in pixels. |
| A parameter appears to have no effect | Sample parameters are specific to each sample. A key that a given sample does not define is accepted and stored without being read, so check the parameter list of the sample being loaded. |
| A wide scan sees the same structure twice | Twelve of the thirteen samples repeat with a 20 µm period. A survey grid whose full extent exceeds that will revisit ground it has already covered. Keep the extent of a scan inside 20 µm, or use the synthetic-feature testbed, which does not repeat. |

One distinction underlies several of these. Settings that belong to the microscope — depth of field, defocus, stage position and tilt — are identical for all thirteen specimens. Parameters that belong to a specimen — how many features it contains, how they are sized, how they are arranged through its thickness — differ from sample to sample, as they must. A knob available on one sample does not necessarily exist on another, and where it does not, it is accepted and ignored rather than raising an error.

# 10. The notebooks

Each writes the same specimen package, server and client, byte for byte; they differ only in what they demonstrate.

| Notebook | Contents |
| --- | --- |
| STEM_Digital_Twin_Kinematical | The twin itself. Specimens, imaging, diffraction, thickness, drift, contamination, depth of field, magnification and stage safety, feature finding, and the backend abstraction. Every demonstration is driven by instrument controls alone. |
| Demo — shape assembly | The synthetic-feature testbed on its own: what its structural parameters do, how far its features are distributed through the foil, and how the width of the tilted in-focus band is set. It closes with an analytic particle projector used as a fidelity reference. |
| Demo — limits and bounds | Live sliders that push each parameter until the image stops making sense, so the working range of each can be read off directly. |
| Appendix A — abTEM multislice | Dynamical diffraction on the same specimen, with frozen phonons, 4D-STEM generation and a py4DSTEM round-trip. |
| Appendix B — EELS | Single-spot spectrum acquisition and the interface a real spectrometer sits behind. |
| Appendix C — Porting to a real instrument | The ten-method specification a backend is built to, with vendor adapter skeletons for ThermoFisher, Nion and JEOL. |
| Appendix D — An ambiguously specified workflow | A survey-and-characterise routine that finds features, ranks them by an explicit policy and drives the stage to the most promising. |

# 11. Moving to a real instrument

A workflow talks to a backend object, never to the twin directly:

def survey_then_focus(backend, fov_um=30.0, current_pA=100.0):
    backend.set_beam(current_pA=current_pA, disabled=False)
    backend.set_fov_um(fov_um)
    af = backend.autofocus()
    return af, backend.acquire_image().shape

Nothing in that function knows what it is driving. Deploying means constructing a different backend and passing it in. The interface is ten methods:

| Method | Contract |
| --- | --- |
| get_stage() / set_stage(x, y, z, a, b, relative) | x, y, z in metres; a, b in degrees |
| get_beam() / set_beam(current_pA, voltage_kV, disabled) | pA and kV; disabled defaults to true |
| set_mode(mode) | "IMG" or "DIFF" |
| set_fov_um(fov_um) | field of view in micrometres |
| get_magnification() / set_magnification(magnification) | returns both magnification and field of view |
| acquire_image() | a plain two-dimensional array |
| autofocus() | a dictionary containing a converged flag |

Adapters for ThermoFisher (AutoScript), Nion and JEOL (PyJEM) are provided as skeletons carrying the SDK calls as documented. They are unvalidated on hardware: axis signs, unit conventions, settle behaviour and detector identity have to be checked at the instrument by a person.

We would not recommend generating an adapter automatically. None of those facts can be established from documentation, and code that is wrong in any of them is not distinguishable from correct code by inspection — a drift correction that runs the wrong way, or a blanking call that is accepted and ignored, is expressed on hardware rather than caught in review. What the interface guarantees is the shape of the call and that a workflow written to it carries no simulation-specific dependency; verifying that a given instrument satisfies it remains a manual step.